RamanSense: An Adaptive Algorithm for Enhanced SERS Analysis and Precise Detection of Low-Concentration Molecules in Complex Media.

Abstract

Detecting low-concentration molecules in complex media presents a significant challenge in Raman spectroscopy, particularly due to overlapping peaks. This study introduces RamanSense, a robust semi-automated algorithm designed to enhance the analysis of Surface-Enhanced Raman Spectroscopy (SERS) data. The algorithm incorporates advanced preprocessing techniques, including region-of-interest (ROI) selection, noise reduction, baseline subtraction, and Min-Max normalization, which ensure the preservation of critical spectral features. The performance of RamanSense was validated using 4-Aminothiophenol (4-ATP) as a model molecule at a 1 μM concentration, achieving a Signal-to-Noise Ratio (SNR) < 2% and precise spectral peak detection, even in complex media. Additionally, the Spectrum Profile (SP) indicator enables comparative spectral analysis, enhancing the algorithm’s ability to differentiate molecular species and variations. RamanSense’s adaptability, achieved through customizable control parameters, provides superior performance over conventional methods such as PCA and PLS by retaining more spectral detail. This work represents a significant advancement in Raman spectroscopy, enabling precise detection and analysis of molecular signatures with applications in biomedical diagnostics, material science, and beyond.

Introduction

Raman Spectroscopy is a well-established analytical technique that probes molecular structures through the inelastic scattering of photons. By measuring shifts in energy levels resulting from molecular vibrations, it reveals detailed information about a sample's composition. However, while effective for high-concentration samples, Raman spectroscopy struggles with sensitivity when analyzing low-concentration molecules, particularly in complex media. Overlapping peaks, noise, and baseline interference further compound these challenges. These limitations have driven the development of advanced techniques like Surface-Enhanced Raman Spectroscopy (SERS) and Tip-Enhanced Raman Spectroscopy (TERS). SERS, by utilizing nanostructured surfaces, significantly amplifies Raman signals, enabling trace-level detection with high reproducibility and minimal pre-analytical requirements. Similarly, TERS enhances spatial resolution through nanoscale tip interactions, providing sub-diffraction-level morphological and chemical insights.

Despite these advancements, challenges persist. SERS is hindered by substrate reproducibility and standardization issues, while TERS demands precise tip-sample alignment, making it technically complex. Furthermore, traditional data analysis methods, such as Principal Component Analysis (PCA) and Hierarchical Cluster Analysis (HCA), though effective for dimensionality reduction and pattern recognition, often neglect critical spectral details. PCA, for instance, is sensitive to noise and outliers, while risking information loss during dimensionality reduction. These challenges underscore the need for innovative algorithms that maintain spectral integrity, enhance sensitivity, and address the limitations of conventional techniques.

In this context, we introduce RamanSense, a robust semi-automated algorithm designed to enhance the analysis of Raman spectra, particularly for identifying low-concentration molecular signatures in complex media. RamanSense incorporates adaptive preprocessing steps, including region-of-interest (ROI) selection, noise reduction, baseline subtraction, Min-Max normalization, and Spectrum Profile (SP) comparison. These steps address critical challenges such as overlapping peaks, noise interference, and baseline drift, ensuring the preservation of critical spectral features. Unlike fully automated methods, RamanSense offers customizable control parameters, including thresholds and window sizes, enabling fine-tuned spectral analysis tailored to diverse datasets.

The significance of this work lies in its ability to bridge critical gaps in Raman spectral analysis. By improving sensitivity, specificity, and reproducibility, RamanSense holds significant potential for applications such as biomarker detection in biomedical diagnostics, environmental monitoring, and material science. Previous studies have demonstrated the potential of advanced SERS methodologies—Matschulat et al. (2010) developed multiplexed SERS imaging for biological systems, Koo et al. (2018) clinically validated SERS for cancer risk prediction, and Premachandran et al. (2022) showcased non-invasive liquid biopsies using 3D nanoprobes. However, these studies often emphasize experimental advancements over robust data analysis techniques. Our work complements such methodologies by providing a versatile analytical tool capable of improving reliability and reproducibility in complex spectral environments.

Validation of RamanSense was performed using 4-Aminothiophenol (4-ATP) as a model molecule. Results demonstrated an SNR of 1.48–1.62% across thresholds, successfully differentiating low-concentration signals from background noise and baseline interference. The algorithm also maintained spectral integrity while detecting 4-ATP at concentrations as low as 1 μM in complex media. Additionally, the Spectrum Profile (SP) indicator facilitates comparative spectral analysis, further enhancing its adaptability across different conditions. These findings underscore RamanSense's value for advancing Raman spectral analysis, enabling precise detection of low-concentration molecules in challenging environments.

Results and discussion

To identify the Surface-enhanced Raman Spectroscopy (SERS) spectrum signatures for low-concentration molecules (beyond the detection limit of conventional Raman) in complex media, we developed RamanSense, a detailed pre-processing Algorithm to maximize the SNR to identify the molecule peaks precisely through pattern recognition. Here, we used 4-Aminothiophenol (4-ATP (molecule at a 1μM concentration as our model molecule. Though, the algorithm can be adaptable to various molecules at different medias. The process included a series of steps, beginning with converting the raw SERS data into numerical columns for the Raman shift (the x-axis) and intensity values (the y-axis). Figures S.3 and S.4 illustrate this process after selecting the region of interest for the molecule's signature (in our case 1000≤Xn≤1650 for 4-ATP molecule).

The region of interest (ROI) was determined as the following:

| Dn=1 0if Xmin≤Xn≤Xmax otherwise | (1) |
| --- | --- |

, where Dn values are binary weight of either 1 or 0 accordingly, and Xn is the Raman shift (x-axis).  The Dn=1 and Dn= 0 represents the ROI and non interested region, respectively.

To define the  Xmin  and Xmax for Dn = 1, we performed the first derivative of the intensity function across a broader range (in our case, over 50 points). In case the first derivative is equal or approximated to zero, means Dn = 0, i.e., non-interested region, otherwise when the first derivative is equal to a value that is greater than 0 means Dn = 1, i.e., the ROI. We identified regions where the first derivative effectively approached zero:

| dY(n±m)dX(n±m)≈0 | (2) |
| --- | --- |

where n represents the window center. m is an integer that moves along the wavenumber until negative infinity. Yn represents the intensity value from the raw data, with n ranging from -m to m. These regions where Dn = 0 corresponded to areas with a negligible signal contribution. We assumed that the first derivative over a broader range would help identify areas with or without significant peaks, and therefore, having or lacking important data, respectively. Consequently, these regions were considered or neglected in our analysis to focus on the relevant spectral signals. This careful selection ensured that only relevant spectral data was retained. Notably, the ROI can vary depending on the molecule, allowing our algorithm to be customized for specific cases. This adaptability offers a significant advantage over methods, such as Principal Component Analysis (PCA), which typically do not involve a targeted selection of the ROI. Following the selection of the ROI, the next step involves applying noise reduction techniques to ensure the data is as clean and accurate as possible. Regions, where the first derivative does not approach zero and at Dn = 1, indicate significant peaks corresponding to molecular signatures that require further analysis. Noise reduction was achieved through a smoothing technique, specifically a moving average with a small sliding window (in our case, a sliding window of 2). This method was applied iteratively over multiple cycles (in our case c=15 cycles), to enhance data clarity and accuracy. The moving average was calculated using:

| MAven±m*c=12*m+1*n-mn+mYnc | (3) |
| --- | --- |

Where n is the center of the window, and m is an integer that determines the window width, 2m + 1 and c is the number of iterations. It's important to note that m cannot equal 0 or the maximum possible value. When m=0, the result would merely be the original Yn​ data without any averaging effect. Conversely, when m equals the maximum possible value, all data points would be averaged into a single value, thereby eliminating any meaningful spectral features. We also assumed that the intensities within the wavenumber inter Xminto  Xmax follow a nonlinear Raman spectrum. This nonlinearity may arise due to competing phenomena such as fluorescence emission, which can interfere with the Raman scattering process. Therefore, we selected this averaging technique because it is particularly well-suited to the nonlinear data. Unlike the harmonic average, this method effectively accommodates lower intensities, which tend to exhibit lower errors or standard deviations. Additionally, given the nonlinear nature of the data, a linear filter would not be effective in this case, so we needed to move windows that transition from one interval to another. The interval width remains consistent in every calibration, enabling a purer signal and a real one—figure S.7 illustrates the real spectrum after smoothing. By dividing the moving average into windows of an interval, we ensure that the first derivative of the window interval remains constant, thereby converting the nonlinear function into a linear one.

Various noise reduction techniques were evaluated, including digital smoothing filters, Fourier Transform, and Wavelet Transform methods. Digital smoothing filters, such as the moving average method and window moving binomial smoothing method, are commonly used for reducing random noise and enhancing the signal-to-noise ratio (SNR). While effective, these methods can sometimes obscure important spectral features, particularly in spectra with sharp peaks. The Fourier Transform is proficient at filtering high-frequency noise but is most effective when there is a significant frequency difference between the signal and noise. The Wavelet Transform, which involves multi-resolution frequency analysis, requires precise selection of the wavelet mother function, decomposition layers, and thresholds. Although this method can be effective, it struggles to extract very small Raman signals that are submerged by background noise.

In our study, we selected the moving average method for noise reduction due to its optimal balance between simplicity and effectiveness. This technique, utilizing a sliding window applied iteratively over multiple cycles, was found to smooth the data adequately while preserving crucial spectral features. Compared to more complex techniques such as Savitzky-Golay smoothing or Wavelet Transforms, which require intricate parameter adjustments and are computationally intensive, the moving average method offered a practical solution. Its straightforward implementation enabled effective noise reduction without compromising the integrity of the spectral data.

This method significantly improved signal clarity by reducing noise and enhancing the detection of subtle spectral features. Consequently, our approach improves both the accuracy and reliability of the spectral analysis, especially in comparison to PCA, which is more sensitive to noise and outliers.

Selecting the appropriate window width for smoothing was essential. A wider window yields a smoother function but can also obscure valid data, while a wider window improved the signal-to-noise ratio (SNR), it did so at the cost of some data loss and resulted in broader peaks. (Figure S.4 vs. Figure S.5). Conversely, a narrower window may preserve more data, but it can also leave significant noise intact. (Figure S.4 vs. Figure S.6).

To determine the optimal window width, we pre-treated the raw data with various moving average ranges and tested different windows and different numbers of iterations. Selecting the optimal window was prioritized over merely improving the SNR, as it balances SNR enhancement with effective data preservation, as illustrated in Figure S.7.  This iterative approach not only improved the SNR but also enabled effective deconvolution of overlapping peaks that were previously obscured by noise in the raw data. Consequently, this allowed for more accurate identification of spectral peaks for the species within the complex media. The number of iterations was chosen to maximize SNR enhancement while ensuring that data integrity was maintained. After smoothing, background noise was determined using a statistical technique—the moving median with a small sliding window (in our case, a sliding window of 7). For this case, we used a window size of s=2m + 1=in our case, 2*7+1=15  is odd, with the median calculated as:

| MMedianodd s = n-mn+ms + 12 term | (4) |
| --- | --- |

Where s the number of noised-reduced intensities in the window, and determining the optimal moving median window, which helps to detect and reduce the background noise surrounding the peak area, was crucial. background noise removal can be achieved using methods such as polynomial fitting, Wavelet Transform, or least squares-based techniques. However, we chose the moving median method due to its robustness in handling complex spectra where the Raman signal may be obscured by background noise, such as fluorescence from surrounding media or contaminants. Unlike polynomial fitting, which requires manual adjustment and is sensitive to the profile of the Raman spectrum, or Wavelet Transform techniques that demand precise parameter selection, the moving median method provides a more straightforward and effective approach. It excels in situations where overlapping peaks are present, offering accurate background subtraction without the need for extensive parameter manipulation. The moving median method minimizes the risk of losing critical spectral information, ensuring that the true Raman signal is preserved for reliable analysis. It effectively smooths the data while preserving essential spectral features. Additionally, this method is less sensitive to outliers and can handle the nonlinearities often present in Raman spectra, making it particularly suitable for maintaining the integrity of the spectral data while removing noise.

We observed that smaller window sizes (Figure S.9) yielded more effective background reduction, and relatively low noise levels providing better smoothing and preserving important spectral information such as the peaks' sharpness and shape. In contrast, larger window sizes (Figure S.8) led to more smooth spectra with insufficient background reduction and potential loss of important spectral details.  This careful selection of window size allowed for improved peak detection and preservation of critical spectral information. Following background detection with the moving median, a baseline subtraction method was applied. The baseline was subtracted from the moving average using the following function:

| Hn=IFYMAVG n<>n±m*c-YMMED n<>n±m>0.1,YMAVG n<>n±m*c-YMMED n<>n±m,0 | (5) |
| --- | --- |

Where the IF statement performs a logical test and returns a value based on whether the test is true or false. Hn represents the baseline subtraction function. We applied a positive threshold value (in our case it was set to 0.1) to neutralize noise and eliminate negative values. This method effectively subtracts unwanted background noise, refining the accuracy of peak detection. The moving median was applied to identify peaks by analyzing the response to the baseline subtraction.

This approach proved crucial in distinguishing true spectral peaks from baseline noise, thereby enhancing the precision and reliability of peak identification in the analysis of complex spectral data. Our baseline subtraction method, which leverages the moving median approach, offers distinct advantages over other techniques. Traditional methods like polynomial fitting, discrete wavelet transform, and rubber-band methods involve assumptions about the baseline's origin or the spectral profile. For instance, polynomial fitting assumes a baseline can be approximated by a low-order polynomial, which might not hold true for complex spectra. Similarly, the discrete wavelet transform method decomposes the spectrum into approximation and detail components, but this requires precise tuning and may not be suitable for Raman spectral analysis. The moving median method, on the other hand, defines the baseline as the median of the surrounding datapoints within a window, effectively ignoring peaks and focusing on the background. This method is particularly effective in handling spectra with overlapping peaks and varying baselines, as it does not require extensive parameter tuning. The simplicity and robustness of this approach make it well-suited for our analysis. By applying the moving median method, we minimize the risk of losing critical spectral information, ensuring that the true Raman signal is preserved. This approach is less sensitive to outliers and can effectively handle the nonlinearities present in Raman spectra, making it particularly suitable for maintaining the integrity of spectral data while removing background noise.

Following the baseline subtraction, normalization was applied as the final preprocessing step to eliminate the influence of fluctuations in excitation intensity or changes in focusing. This step ensures that such variations do not distort the spectral analysis. The normalization process involved dividing the baseline-subtracted data by the moving maximum function, with reference to the moving median, the baseline correction was calculated as follows:

For the same moving max and moving median windows:

| In=HnYMMAX n<>n±m-YMMED n<>n±m*100 | (6) |
| --- | --- |

Where:

| MMAXn±m=n-mn+mMAX [Yn] | (7) |
| --- | --- |

The MMAX function returns the maximum value within a specified range of intensities, identifying the peak intensities in the spectrum. The function In calculates the percentage value of the baseline-subtracted signal (Hn ​), normalized to the difference between the moving maximum values and the moving median values (in our case within a window of 7). This normalization process refines the accuracy of the spectral analysis by accounting for variations in experimental conditions, ensuring that peak intensities are consistently represented across the dataset. The method used for normalizing Raman data significantly influences the interpretation of results, as it affects how spectral features are preserved and analyzed. Compared to traditional methods like SNV and region norming, MMAX normalization is better suited for Raman spectra because it preserves the distinct molecular patterns essential for accurate analysis. While SNV centers and scales each spectrum to correct for scatter effects, it can inadvertently distort local peak intensities and molecular patterns. In contrast, MMAX normalization focuses on local peak intensities, ensuring that the integrity of Raman peaks is maintained, leading to more reliable spectral analysis.

After examining results where different upper (MMax) and lower (MMed) window widths led to undesirable distortions in the results, therefore we decided it better to set both the upper and lower window widths to be equal. In this case, we avoided division by zero and receiving an error message when calculating the normalized value and we avoided receiving negative values because in the case of similar windows, the values of the moving max will always be greater than the values of the moving median, and equation 6 is well suited for this case, the denominator will give only the absolute values. Since only positive values appear in the function Hn, negative values will also not appear here. So far, the differences in readings between the noise-reduced data (MAvg) and the moving median depending on the window width have been calculated. Figures S.10 and S.11 demonstrate the impact of different window widths on normalization. The wider the window, the lower the moving median values and the more readings appear positive in the peak area. The MAvg function was then displayed in the graph as a normalized function using Min-max scaling, calculated as follows:

| NA%=YMAVG n<>n±m*c-YMinYMax-YMin*100 | (8) |
| --- | --- |

Where YMAVG n<>n±m*c is the moving averge intensity over c iterations, Ymax is the maximum intensity value in the MAvg function which yields the highest peak intensity and  Ymin is the minimum intensity value in the graph not necessarily a peak’s intensity.

This normalization method scales the moving average intensity to a percentage, allowing for clearer comparisons and reducing the impact of intensity variations across the dataset. Min-max scaling was employed to normalize the intensity values, ensuring that the entire range of data is proportionally adjusted. While other normalization methods like z-score scaling can normalize the variance of each variable to 1, equalizing the importance of all variables, this approach is not recommended for spectral data such as Raman spectra. Z-score scaling may obscure or distort spectral patterns that are critical for further analysis, making min-max scaling a more appropriate choice for maintaining the integrity of the spectral features.

In Figure S.12, The algorithm model function is displayed in the graph as a clustered column style to emphasize peak detection. After preprocessing the data, we began modelling to identify the real peaks that related to the molecule patterns and to identify the background noise, for this goal, the algorithm enhanced graph resolution and reduced the width noise of the model. This involved performing an integral function on the baseline correction values to calculate the cumulative sum of function I from n=1 to n. Since only positive values appear in function I, the value in function J for a given "n" will always be greater than or equal to the value at "n-1":

| Jn=i=1n(I)i | (9) |
| --- | --- |

We then calculated the positive first derivative function K to isolate the peaks as follows:

| Kn=IF(Jn-Jn-1>0, Jn-Jn-1,0) | (10) |
| --- | --- |

In general, only values exceeding a specified threshold are displayed in the function. Larger threshold values result in fewer normalized values participating in the display. For instance, setting the threshold to 0 will include all normalized values, while setting it to 99.99 will exclude all normalized values except those near 100, showing only the single readings at the precise location of the peaks. (Figure S.12 vs Figure S.13). Accordingly, a threshold of 100 isn't recommended due to no single value will appear in the graph. We set the threshold to 0 when the targeted molecule exists in the same media and set it to 99.99 for accurate and precise peak detection when there are different medias where the targeted molecule exists, also to detect peak shifts and their relation to different medias.

This method advances peak identification by focusing on the most significant values. We then applied a step function U recording a value of 1 only for records where a positive value exists to distinguish significant peaks. Since the functions can contain values greater than 1, we use the PRODUCT function to ensure a result displays either 0 or 1:

| Un= IF(PRODUCTHn, Kn, Dn=0, 0, INTPRODUCTHn, Kn, DnPRODUCTHn, Kn, Dn) | (11) |
| --- | --- |

The INT function rounds numbers down to the nearest integer. This formula multiplies the values in the range together, yielding a convolution technique, PRODUCT that multiplies three functions: the H baseline subtraction, the K positive first derivative, and D the step function of 1 or 0 for the ROI.  If the product is greater than 0, then Un= 1 represents relevant peaks otherwise Un = 0  represents the non-relevant peaks. We use the formula Un as a term to determine the existence of the real peaks inside the ROI, relevant peaks of the targeted molecule, and also determine the real peaks of the media inside the ROI. Finally, we used the SUMPRODUCT function to calculate Vn, which represents the graph of the algorithm model:

| Vn=SUMPRODUCT(Hn:In, 0:1) *MAX (Un, 0) | (12) |
| --- | --- |

The SUMPRODUCT function multiplies corresponding components in the given arrays and returns the sum of those products. This method enabled us to understand the peaks inside the media and to conclude the existing of the targeted molecule or other molecules through pattern recognition. The step function U used as a condition for displaying the relevant values of function I, enables the displaying of the spectral peaks that represent the signature of the given molecule in this case 4-ATP.

The algorithm was designed as a semi-automated analysis, with five control parameters to optimize its performance for the specific dataset. These parameters include (ⅰ) the number of iterations, (in our case, set to 15 cycles); (ⅱ) the sliding window size for the moving average, (in our case, set to 2); (ⅲ) the normalization window for the moving maximum, aligned with the moving median, (in our case, window set to 7); (ⅳ) a positive threshold value (in our case of 0.1) to mitigate noise and eliminate negative values from the baseline subtraction; and (ⅴ) a threshold value, adjustable to 0 or 99.99 in our case, to selectively display normalized values as clustered columns in the graph. These customizable controls enhance the algorithm’s adaptability and precision in spectral analysis, also enable us to pause the process in each step to understand the effect of every step of the process and analysis it in order to adjust it accordingly.

In contrast to conventional methods like Principal Component Analysis (PCA), Partial Least Squares (PLS), Principal Component Regression (PCR), and Support Vector Machines (SVM), which are commonly used for general pattern recognition and data simplification in spectral analysis, our approach offers enhanced specificity and precision in peak detection for Raman spectroscopy. While PCA, PLS, PCR, and SVM effectively manage high-dimensional data and identify broad trends, they often compromise the resolution of localized spectral features by focusing on overall data variance or classification boundaries. Specifically, PCA and PCR aim to capture the maximum variance in the data, which can sometimes lead to the loss of critical peak details. PLS, though designed to model relationships between spectral data and response variables, may still smooth out fine spectral nuances due to its emphasis on maximizing covariance between predictors and responses rather than individual peak features. Similarly, SVM, while powerful for classification tasks, does not inherently prioritize the preservation of intricate spectral features essential for detailed analysis. Our algorithm, however, is designed to retain these subtle yet crucial details, ensuring a more accurate and reliable analysis of Raman spectra.

On the other hand, our model, utilizing Equations 9 through 12, integrates customizable parameters for baseline correction, peak normalization, and threshold adjustment, ensuring precise peak detection while preserving critical molecular signatures. The approach involves cumulative sum for enhancing signal detection (Equation 9), first derivative for isolating peaks (Equation 10), step function for distinguishing significant peaks (Equation 11), and integration for accurate peak representation (Equation 12). This method maintains the integrity of important spectral features, offering a more refined analysis that is particularly advantageous for applications requiring detailed molecular pattern recognition, where traditional methods might compromise the resolution of localized spectral data.

Calibrations

To ensure data accuracy and reliability, we performed multiple calibrations, emphasizing critical metrics including the spectrum profile (SP), standard deviation (SD), signal-to-noise ratio (SNR), variability of the peaks (VAR), and peak number detection. Starting with the spectrum profile which serves as an indicator to compare different spectra (or spots) for the same molecule. It is calculated as follows:

| Mn= Xn*Vn | (13) |
| --- | --- |

And

| SP=i=1nMii=1nUi | (14) |
| --- | --- |

Where Mn is the multiplication of the intensities (Peaks) in function V multiplied by the wavenumber (Raman shift). Xn represent the wavenumber values from the raw data. Vn  the algorithm model values. U is the step function for peak detection. The (SP) method offers a significant advantage over other methods like Green's function, subtracted deconvolution, and similarity functions, particularly in applications requiring precise spectral comparison. Unlike Green's function, which focuses on theoretical modeling of molecular responses, the SP method provides a practical comparative metric that accounts for both intensity and wavenumber variations, making it more effective for real-world spectral analysis. Compared to subtracted deconvolution, which enhances spectral resolution by removing broadening effects but can introduce significant standard deviation in noisy regions between peaks, the SP method not only preserves but also highlights critical peak details. Through its integration of peak detection mechanisms, the SP method provides a more accurate and comprehensive tool for spectral analysis, reducing the risk of error in noise exposure areas and ensuring the integrity of the spectral data. While similarity functions offer a straightforward approach for general spectral matching, they lack the specificity and customization of the SP method, which can be finely tuned to detect and compare subtle spectral differences, particularly in periodic results, making it ideal for applications that demand detailed molecular pattern recognition. Additionally, we calculated the standard deviation and average of spectrum profiles, (in our case from 10 different spots) using the following formulas:

ⅰ. Standard Deviation (STDEVA):

| STDEVAx1,x2,…, xn=i=1n(xi-x)2n-1 ⅱ. Average (AVERAGE): | (15) |
| --- | --- |
| AVERAGEx1,x2,…, xn=i=1nxin | (16) |

Where n is the number of profiles which equal (in our case to 10), xi represents the SPOT profile value at spectrum number i and x represents the average of all (in our case 10) spot profiles.

We then calculated the signal-to-noise ratio (SNR) using:

| SNR=STDEVASPOT1:SPOT10AVERAGE SPOT1:SPOT10=NoiseSignal | (17) |
| --- | --- |

Where we defined it to be noise as a percentage of the signal by representing the noise as the standard deviation of the (in our case 10) spectrum profiles and the signal as the average of the (in our case 10) spectrum profiles. where the standard deviation measures the variability of intensity value within the spectrum, considered as noise, While the average measures the signal of interest within the spectrums. When the SNR approaches lower values, it indicates that the signal is higher relative to the noise, reflecting better accuracy of the sensor and data quality. Calculating the signal-to-noise ratio helps evaluate the quality of the spectra by comparing the strength of relevant spectral features (signal) to the background noise level. A lower SNR indicates clearer and more reliable spectral data, making it easier to identify and analyze specific components of interest. (In our case the SNR was 1.62%) with the fifth control parameter set at a threshold value of 0, and 1.48% when set at 99.99. This demonstrates that improved peak identification results in enhanced SNR. Compared to traditional SNR approaches, such as those used in Amplitude Modulation (AM) and Frequency Modulation (FM) systems, are ineffective for Raman spectroscopy because they assume specific noise models and modulation schemes that do not align with the complex and varying noise characteristics found in Raman spectral data. Other SNR calculations typically address baseline increases and random fluctuations. The first method evaluates SNR by comparing the Raman peak intensity to both baseline and signal noise, offering insight into overall detector performance, particularly against extrinsic noise sources like CCD interference. The second method focuses on post-baseline subtraction, assessing SNR by dividing peak signal intensity by the standard deviation of random fluctuations, widely used to ensure reliable spectral detection. Our approach, however, broadens the scope by directly comparing the variability within spectral data (standard deviation) to the mean signal strength, thereby providing a more comprehensive evaluation of data quality across multiple samples. This method enhances reliability, particularly in applications requiring precise spectral comparisons, where traditional methods might overlook subtle but critical variations in spectral data.

The final calibration was the variability of the peaks, which quantifies how much the intensity values deviate from their average across the spectrum. Higher variance values indicate greater variability, suggesting a diverse distribution of intensity peaks. Conversely, lower variance values suggest a more uniform intensity distribution, with peaks clustered closer to the average intensity. This analysis aids in understanding the distribution pattern of spectral intensities. The variability of the peaks was calculated using the formula:

| VAR.Px1,x2,…, xn=i=1n(xi-x)2n | (18) |
| --- | --- |

Where n is the number of profiles which equal (in our case to 10), xi represents the SPOT profile value at spectrum number i and x represents the average of all (in our case 10) spot profiles. In the analysis of spectral data, different methods for calculating peak variability include standard deviation, coefficient of variation (CV), root mean square error (RMSE), mean absolute deviation (MAD), and interquartile range (IQR). Standard deviation and CV are commonly used but can be sensitive to outliers and may not fully capture peak-specific variations. RMSE focuses on prediction accuracy rather than inherent peak variability, making it less relevant for spectral analysis. MAD, while robust to outliers, does not directly address peak-to-peak differences. IQR measures variability within the central portion of the data but may not be applicable to peak-specific analysis. In contrast, our method calculates variability based on peak intensity deviations from the average, providing a more precise assessment of peak characteristics and distribution. This approach effectively manages the complexity of spectral data, offering enhanced precision for analyzing intricate patterns and variations in peak analysis. Lastly, the peak detection formula is defined as follows, the Local Maxima Method:

| Ln=IFANDKn>Kn-1,Kn>Kn+1,1,0 | (19) |
| --- | --- |

where L represents the peak counting function, which assigns a value of 1 exclusively to data points that constitute a significant peak. Specifically, a peak is identified when the reading at Kn​ is greater than both the preceding reading Kn-1 and the subsequent reading Kn-1​. The total number of peaks is then calculated by summing all the values assigned by the function L. Additionally, the condition that the function U also contains the value 1 serves as a weight function, further refining the peak detection process. In spectral peak detection, several methods are commonly used, each with its strengths and limitations. The Derivative-Based Method uses changes in the derivative to detect peaks, which may miss smaller or broad peaks. Peak Fitting involves modeling peaks with functions like Gaussian, offering high precision but requiring accurate parameterization. The Wavelet Transform provides multi-resolution analysis, capturing peaks at various scales but can be complex to implement. In contrast, our method involves a multi-step process that includes normalized baseline subtraction, integral calculation, positive first derivative with a threshold, and the Local Maxima Method for peak detection. This approach is tailored for high precision in complex spectral data by addressing baseline noise, amplifying significant features, and focusing on local maxima to accurately identify peaks. The detailed pre-processing steps and the model improve the specificity of peak detection, making our method particularly effective for complex spectra with overlapping or subtle peaks.

In summary, the algorithm helps identify real peaks and differentiate them from fake ones, examining changes in peak intensity and peak position (Raman shift).

Model evaluation

The algorithm effectively identifies low-concentration molecules within complex media. In tests with and without 4-ATP, it accurately detected 4-ATP even at low concentrations (1μM), demonstrating its robustness. Preprocessing steps such as baseline subtraction and normalization, along with advanced peak detection techniques, ensure precise and reliable results. Figures S.14a-e illustrate the algorithm's capability to discern low-concentration molecular signatures inside complex mixtures.

Conclusion

The development and implementation of an advanced algorithm for enhancing the analysis of SERS spectra has proven highly effective in identifying spectral signatures of low-concentration molecules within complex media. Utilizing 4-Aminothiophenol (4-ATP) as our model molecule, the algorithm successfully transformed raw spectral data into a format suitable for detailed and precise analysis. Our approach included a series of preprocessing steps such as noise reduction, background subtraction, normalization, and peak detection. These steps were instrumental in significantly improving the clarity and accuracy of the spectral data. Techniques like moving average, moving median, and baseline subtraction were essential in reducing noise and enhancing signal integrity. Meanwhile, our normalization and peak identification methods enabled the clear recognition of distinct spectral features.

A notable feature of the algorithm is its ability to distinguish between real and false peaks, examining changes in peak intensity and position (Raman shift). This capability is essential for accurately detecting and analyzing low-concentration molecules within complex samples. Additionally, our iterative approach to optimizing window widths for smoothing and baseline correction ensured the preservation of critical spectral details while minimizing noise. However, the current algorithm is not without limitations. Its performance is heavily dependent on the choice of control parameters, which may require manual adjustment and optimization for different datasets. Furthermore, processing large datasets can be time-consuming, which could restrict its use in time-sensitive scenarios.

To address these limitations and enhance our research, future efforts will focus on several key areas. One major objective is to automate parameter selection using machine learning techniques, enabling adaptive optimization of the algorithm for various types of spectral data. We also plan to integrate our algorithm with Principal Component Analysis (PCA). This combined approach aims to improve the identification of low-concentration signatures, specifically targeting mi17RNA in blood samples. Such integration is expected to yield a powerful tool for biomedical diagnostics applications, facilitating the detection of critical biomarkers at low concentrations with high accuracy. Improving the computational efficiency of the algorithm will also be crucial for handling large datasets more effectively, making it suitable for high-throughput analysis and real-time applications. Additionally, we aim to expand the algorithm's applicability to a broader range of molecules and complex media, enhancing its versatility in SERS spectrum analysis.

In summary, our algorithm represents a significant advancement in the field of SERS spectra analysis by accurately identifying and analyzing low-concentration molecules within complex samples. This capability is vital for further advancements in various scientific and biomedical fields, offering a reliable tool for detecting critical biomarkers and enhancing spectral data analysis.